Cerberus Simulator System Definition
2020 April 14

Note: all pressure values are gauge readings referenced to 1 atm = 0
X[i] indicated patient specific parameter X for each of two patients on the system

# 1. INPUTS
## 1.1. Patient characteristics (for each of two patients)
1.1.1. Height (H[i]): 110-200 cm
1.1.2. Weight (W[i]): 40-120 kg
1.1.3. Lung Compliance (Cr[i]): 10-70 ml/cmH20 (intervals of 5)
1.1.4. Ratio of arterial partial pressure of oxygen to fraction of inspired oxygen (PF[i]): range 100 (severe dysfunction) - 500 (Normal)
For simplicity we are combining a whole bunch of physiological parameters into one.
## 1.2. Ventilator settings
### 1. 2.1. Primary ventilator settings
1.2.1.1. Pressure Control Level (PC) : 0-70 cmH20 (intervals of 5)
1.2.1.2. Positive End-Expiratory Pressure Level (PEEP0): Should always be set to 0 and somehow *gently* highlighted. Setting this above zero is a common error.
1.2.1.3. Respiratory Rate (RR): 2-40 breaths/minute
1.2.1.4. Inspiratory Time (IT): 0.5-(0.6/RR) s (That is inspiration should be no more than 60% of the respiratory cycle)
1.2.1.5. Inspiratory Time:Expiratory Time Ratio (IER): Discrete: 1.5:1 ; 1:1 ; 1:2 ; 1:3  (This represents the same info at IT but in a different form that will be more familiar for some users. For simplicity we can make this adjustable and make IT automatically calculated and displayed next to it. [ I:E = IT:((1/RR)-IT) ]
### 1.2.2. Secondary circuits settings (for each of two secondary circuits
1.2.2.1. Fresh Gas Oxygen Flow (FGFO[i]): 0-15 LPM (intervals of 1 for all flow variables)
1.2.2.2. Fresh Gas Air Flow (FGFA[i]): 0-15LPM
1.2.2.3. PEEP Valve Setting (PEEP[i]): 0-25 cmH20
# 2. Outputs
(values outside normal range should be highlighted)
## 2.1. Patient Arterial Blood Gas Analysis (for each patient)
2.1.1. pH[i]: Normal 7.3-7.4. Function of PaCO2[i] and HCO3[i];
- [H+] nM/L = 24 * PaCO2 (mmHg)/[HCO3] mM/L
- pH = -log10[H+]
2.1.2. Partial pressure of carbon dioxide (PaCO2[i]): N 35-45 mmHg (note different pressure units from above pressure values)
- Function of minute alveolar ventilation (3.1) and CO2 Production (3.2).
- We can assume that expired gas has similar CO2 partial pressure to the arterial gas, so
- MAV[i]*PaCO2[i] = CO2P[i] (unit conversions needed)
2.1.3. Partial pressure of oxygen (PaO2[i]): N > 100 mmHg PO2 = FiO2 * PaO2:FiO2
2.1.4. Bicarbonate concentration (HCO3[i]): N 22-28 mEq/L
## 2.2. Ventilator derived values (for each patient)
2.2.1. Peak Inspiratory pressure (PIP[i]): range 0-40 cm H2O; note: circuit includes a pop-off valve at 40 so pressure can never go above that
2.2.2. Actual PEEP (APEEP[i]): cm H2O
- Total fresh gas flow: FGFTot[i] = FGFO[i] + FGFA[i] (air + oxygen)
- Intrinsic PEEP (from FGF): PEEPInt[i] = 0 if FGFtot[i] < 7LPM; 3 + (FGFtot[i] – 7)*0.6
- APEEP[i] = PEEP0 + PEEPInt[i] + PEEP[i]
2.2.3. Tidal volume (VT[i]): 50-700 ml  (Normal range depends on weight: 4-8 ml/kg) = 0.5x PIP x Cr for each patient (mean inspiratory pressure assumed to be half of PIP)
2.2.4. Fraction Inspired Oxygen (FiO2[i]): (FGFA*0.21 + FGFO) / (FGFA + FGFO)
# 3. Hidden Variables
## 3.1 Minute alveolar ventilation
3.1.1 Anatomical Dead Space ADS[i] = W[i]/0.45 ml
3.1.2 Secondary Circuit compliance CC[i] = constat CC for simplicity
3.1.3 Alveolar Tidal Volume AVT[i] = VT[i] – ( PIP[i] / 2*CC ) - ADS[i]
approx. assumes linear increase in pressure
3.1.4 Minute Alveolar Ventilation MAV[i] = RR*AVT[i]
3.2 CO2 production rate (VCO2[i]) = 0.4 + (Weight kg / 70 kg ) * 0.2 g / min  (I more or less made this up; average value for 70kg adult is around 0.6 g/min and it varies with weight; for our purposes and the range of adult weights, we are considering that equation is probably good enough
# References
Kassirer JP, Bleich HL. Rapid Estimation of Plasma Carbon Dioxide Tension from pH and Total Carbon Dioxide Content. New England Journal of Medicine. 1965;272: 1067–1068. doi:10.1056/NEJM196505202722007